## Slide 1
EWPAA 2024
Dresden, Germany, September 17-19, 2024
Photocathode characterization using a UV-tunable ultrashort-pulse radiation source
W. Andreas Schroeder
Louis Angeloni and Ir-Jene Shan
Physics Department, University of Illinois Chicago
Theodore Vecchione
Stanford Linear Accelerator Consortium (SLAC)
Department of Energy
DE-SC0020387
UIC

## Slide 2
Outline
Part I: Photocathode characterization
system at UIC
 Design considerations
 Key components:
Tunable UV radiation source
DC gun
Electron (beam) detection
 Calibration and performance
Part II: UV laser cleaning (low power)
 Video!
 QE(t) & MTE(t)   reduction …
 : Experiment vs. Theory

## Slide 3
Design Criteria
Tunable UV radiation source
 Ultrashort pulse laser system PLUS nonlinear optics
Simple 10-20kV DC gun design (+ drift region)
 Straightforward electron trajectory simulation
Sensitive calibrated electron detection
 Dual-plate MCP / P-43 phosphor + CMOS digital camera
Good optical spatial beam stability
 Pinhole / Fourier plane spatial filter (‘top hat’ and near Gaussian beams)
Eliminate intra-pulse space-charge effects
 Radiation source repetition rate > 10MHz
Enhance electron emission
 60 AOI for p-polarized UV radiation to reduce reflection losses

## Slide 4
Photocathode characterization system
 Laser / NLO based sub-picosecond tunable UV radiation source: 3.0 – 5.3eV (235-410nm)
20 W Yb:fiber laser
400 fs, 16.7 MHz, 1030 nm
Nonlinear fiber
(SC-3.7-975)
Detector:
MCP & CMOS camera;
Faraday cup
Non-collinear
sum frequency
UV generation
ΔpT measurement
LBO
LBO
LBO
LBO
LBO
6 W 515 nm
BBO
15 W
0.5 W
SHG
SHG
THG
F
OPA 2
OPA 1
-10 kV
3
s,i
Continuum
(700-950 nm)
Tunable UV (0.01-0.1mW)
2

## Slide 5
Optical parametric amplification
Signal wave amplification (721-946 nm)
 ħωUV = 4.92 - 5.33 eV (3ω SFG)
= 3.72 - 4.13 eV (2ω SFG)
Idler wave amplification (1.13-1.8 µm )
 ħωUV = 4.3 - 4.71 eV (3ω SFG)
= 3.1 - 3.51 eV (2ω SFG)
PLUS: 4ħωlaser = 4.81 eV
3ħωlaser = 3.61 eV
3.1  5.33 eV
UV radiation
(10-100 µW, ~0.5 ps)
Continuum
OPA 1
OPA 2
Amplification range
 OPA 1 and OPA 2 are temperature tuned LBO crystals (NCPM)
Wavelength (nm)
Intensity (a.u.)

## Slide 6
Experiment: UV radiation delivery
 Laser / NLO based tunable UV radiation source: 3.0 – 5.3eV (235-410nm)
20 W Yb:fiber laser
400 fs, 16.7 MHz, 1030 nm
Nonlinear fiber
(SC-3.7-975)
Detector:
MCP & CCD camera;
Faraday cup
Non-collinear
sum frequency
UV generation
ΔpT measurement
LBO
LBO
LBO
LBO
LBO
6 W 515 nm
BBO
15 W
0.5 W
SHG
SHG
THG
F
OPA 2
OPA 1
-10 kV
3
s,i
Continuum
(700-950 nm)
Tunable UV (0.1-1mW)
2
Tunable UV (0.01-0.1mW)
100-200μm aperture
Detector:
MCP & CMOS camera;
Faraday cup
ΔpT measurement
-10 kV
f = 50cm
f = 38cm
OAP
/2

## Slide 7
DC photoelectron gun design
10-20kV DC gun based on design of Togawa et al.
 Simple, proven, and readily fabricated design
 Minimal anode lensing
Material: Polished 316 SS
d = 1cm  EDC = 1-2MV/m
Gun geometry yields uncomplicated E(r,,z) evaluation
 Straightforward photoelectron
trajectory simulation
K. Togawa et al., Phys. Rev. ST Accel. Beams 10 (2007) 020703
Photocathode
Cathode
(-10kV)
Anode
(0V)
e-
6mm 

## Slide 8
MTE calibration
DC gun
Drift region
Detector screen
1cm
42cm
 Electron trajectory simulations using E(r,,z)
Electron trajectory simulation results:
Minimal (and well understood) negative anode lensing
Negligible beam distortion within 0.5mm radius of DC gun axis
(c.f. laser beam size ~100μm)
On detector, near gun axis at 10kV;
pT calibration = 0.445
Trajectory simulations in agreement with Analytical Gaussian (AG) electron pulse propagation model
J.A. Berger and W.A. Schroeder, J. Appl. Phys. 108 (2010) 124905
Photocathode

## Slide 9
Electron beam detection
BOS-18 beam imaging system
 Dual-chevron 10m-pore MCPs: 102  103 gain/plate
(voltage-controllable gain)
 P-43 phosphor screen: 25μm PSF @ 2keV
 Intrinsic MTE resolution of  1meV
for 10keV electrons
2.4m-pixel CMOS digital camera
(FL-20BW; Axiom Optics)
 >80% peak QE @ P-43 phosphor emission
 f1:f2 = 5:8 magnification reduction to camera
 Overall PSF  20m (around 8 pixels)
https://beamimaging.com/product/bos-18/

## Slide 10
QE calibration
Uses ‘canonical’ Rh(110) photocathode
 Theoretical QE dependence
QE = A(ħω)2.89
consistent with measurements
Absolute calibration using Faraday cup
PLUS:
Voltage/gain calibration of MCPs
 QE detection floor of ~10-10
G. Adhikari, P. Riley, and W.A. Schroeder, AIP Advances 9 (2019) 065305
G.W. Gobeli and F.G. Allen, Phys. Rev. 137 (1965) A245-A254

## Slide 11
UV laser cleaning: Rh(110)
Laser parameters:
 = 257nm (ħω = 4.81eV)
PUV  1W/cm2
16.7MHz, 250fs pulses
 Ppeak  200kW/cm2
 No surface melting
(T < 10K)
20-minute video in 25 seconds

## Slide 12
Laser cleaning: Synopsis
Is laser cleaning of photocathode surface simply reducing ?
 Seven selected frames from video
t = 0s
t = 10s
t = 900s
t = 30s
t = 5s
t = 360s
t = 1146s
Rapid changes for t < 30s
Slow changes for t > 60s

## Slide 13
Laser cleaning: MTE(t)
 MTE(t) for Rh(110)
fit by double exponential:
 UV illumination decreases 
(surface roughness constant)
(i) Initial fast  reduction:
 Removal of surface adsorbates
(ii) Subsequent slower  reduction:
 Removal of bonded OH, CO, H …
BOTH surface effects
1 = 0.101min (6s)
2 = 1.133min (68s)

## Slide 14
Laser cleaning: QE
For Rh(110):
and
 with n = 2.89
… only consistent for t < 2 minutes
QE increase after t  2mins. at relatively ‘constant’ MTE
 A bulk effect
 Expulsion of H from photocathode?
 …
QE (a.u.)

## Slide 15
: Theory vs. Experiment
Some convergence between theory and experiment observed
BUT more work to be done … Cu(001), Ta(111), Cr(111), Re(0001), Hf(0001)
 Experimental ‘laser-cleaned’ QE1/n intercept values for ; DFT-based ‘half thin-slab’ calculations (Amsterdam Suite)
Photocathode | expt. (eV) (0.1eV) | DFT (eV) (0.1eV) | PE (eV) | Recommended (eV)
Ag(001) Mo(001) Nb(001) Rh(110) Ta(001) W(001) W(111) | 3.94 3.89 3.94 4.13 3.97 4.08 4.18 | 4.25 4.30 3.52 4.37 3.96 4.02 3.82 | 4.22 (0.04) 4.53 (0.02) 3.97 (0.02) | 4.35 (0.05) 4.46 (0.11) 4.08 (0.17) 4.86 (0.21) 4.10 (0.25) 4.70 (0.06) 4.44 (0.03)
G.N. Derry et al., J. Vac. Sci. & Technol. A 33 (2015) 060801
Preliminary data

## Slide 16
Summary
Part I: Photocathode characterization system at UIC
 Design considerations
 Key components:
Tunable UV radiation source
DC gun
Electron (beam) detection
 Calibration and performance: QE ~ 10-10 and MTE  1-2meV
Part II: UV laser cleaning (low power)
 Video: Increase with t of both QE and MTE
 QE(t) & MTE(t)   reduction (t < 2mins.)
PLUS: Possible bulk effect
  determination: Experiment vs. Theory
… critical for study of near-threshold PE effects

## Slide 17
UIC
Thank you!