## Slide 1: Booster optical beamline
Terry Atkinson, Ji-G. Hwang, S. Wiese
DEELS
2022

## Slide 2: Design Criteria
Transport line as short as possible:
  move output port from sector 5 to sector 16
  extend working area for optical platform
Use only plain mirrors
Simple telescope for slight convergence/collimation
Vacuum window at booster – not beamline
  Wedged window to reduce the transmission of reflections
No down-time regarding source point imaging – high level diagnostic in control room
Installation over 4 stages:
  SD20 sector swap in booster!
  SD21 labyrinth, transport line, IVP, optical table
  SD22 telescope, mechanical upgrades, streak camera
  SD22/23 THz detector, 2. wedged window, mechanical upgrades
Motivation:
Measure the Bunch length over the ramp

## Slide 3
Design consideration 2020
Refractive
optics
2034.71 mm
1001.2 mm
Critical wavelength
shorter than 600 nm
θc ~ 6 mrad
Ec ~ 90 MeV
θ ~ 1 mrad
Spot size without optical components:
σx,y ~ 1 mrad (natural angle) x 11 m (distance from source to last mirror) ~ 11 mm
with the window (Ø 28.7 mm), high losses during transport of the light from low energy beam…
Ji-G. Hwang

## Slide 4
light intensity 2020
Integrated photons over 400 – 800 nm wavelength
@ 2.2 m downstream from source point (first lens)
Numerical simulation by SPECTRA
with opening angle: 11 mrad (r=25 mm)
εx = 100 nm-rad
σt = 50 ps
σE/E0 = 0.1%
Single Bunch charge in booster is <0.4 nC
With a transport line of
1) 7 mirrors (98% reflectivity)
2) 2 lens (97% transmission)
3) 2 windows (99% transmission)
Transmission would be approx. 80 %...
Ji-G. Hwang

## Slide 5
Geometrical and chromatic aberrations 2021
Source to mirror
2034.71 mm
Δzf1lens
zf2opt
1001.2 mm
Due to the space limitation, two options are feasible:
1. f1 = 500 mm, Δzf1lens = 100mm, zf2opt = 741.5 mm
2. f1 = 400 mm, Δzf1lens = 200mm, zf2opt = 578 mm
f1 = 400 mm is more sensitive on angular divergence and f2 position, but it is better in terms of chromatic aberrations.
Ji-G. Hwang

## Slide 6
CAD DESIGN – start to end
S. Wiese
Optical table
Labyrinth
Booster ceiling
IMP
Mirrors
Telescope
Booster
Bending
Magnet

## Slide 7: CAD DESIGN – Labyrinth problem
Massive labyrinth requirements from Rad. Safety (block height represents ceiling thickness!)
+20-degree angle to align to optical table (internal wall and fire door in the way!)
Motorized mirrors in labyrinth (can only be opened in shutdown)
Platform floor extended, new safety railing
SD 2021
To drill the hole in the floor it took 6 diamond drill sets and 3 days!
7

## Slide 8: Mechanical Installation
Cable racks cut
Old Ccd camera, mirror and objective lens removed
1st and 2nd Mirror mounted on a frame
Laser aligned from dipole
SD 2021
Old camera box 15 years great service!
8

## Slide 9: Mirrors
Standard HZb flat mirror design with 2-point spring axis
Bench tested for alignment
“home” issues, new design under development
Mirror test lab ongoing…
9

## Slide 10: First Beam Sept 2021
sss
Laser alignment beam from table and photon beam from booster observed at the intermediate viewing port
Motor control wires too thick to be safely connected to the control box!
controllers/motors got too hot – changed max amp allowance in software
Webcam and laptop used to find first beam at platform
Zero hours of dark-time!
10

## Slide 11: Optical table 2021/22
Translation mirrors to bring the light onto the table
16.5cm working height (streak camera relevant)
Basler ccd camera used to align translation mirrors to the working height
FAST DIODE (HamaMAtsu)
Then the fun starts!
11

## Slide 12: Source point imaging
Setup: (without telescope)
Source point to lens ~ 13m
Lens to camera ~ 1m
standard camera + system at BESSY II
Measured:
SigX – 400um, SigY – 250um
12

## Slide 13
Notes: simple cross-check
Abbe limit @600nm
Without optics (only plane mirrors)
Source -> focusing lens 13m
Spot size at focusing lens 20mm (2-inch lens)
Half the angular acceptance 10mm / 13m = 0.77mrad
Limit:
lambda / 2n sin(theta) = 600e-9 / 2*1*sin(0.77) = 243um

## Slide 14: Fast Diode (rise time ~ 50ps)
Setup:
diode -> Bias-T -> two Amps -> scope
trigger FCT5
Bunch length ~ 1/sqrt(V')
Multi-Gigabit Transmission over multimode optical fibre: Stefano Sottacchi p105/106
0.7mA SB current, @20ms in ramp
604kV, avg. St = 26.35ps
714kV, avg. St = 24.22ps
14

## Slide 15
Gaussian Bandwith notes
12GHz
12GHz
rise time = 50 ps # Hamamatsu

## Slide 16
Optical table 2022/23
PC + Monitor
Streak Camera
Fast diode
Camera
Research and development
Scope
90%
source

## Slide 17: Thanks to
G. Schiwietz, M. Ries, G. Klemz, G. Rehm+Jens K.
D. Schüler, H. Glass, B. Kuner
Installation over 4 stages:
  SD20 sector swap
  SD21 labyrinth, transport line, IVP, optical table
  SD22 telescope, mechanical upgrades –> finished last month!
    BSc or Master Student
    Test telescope with beam
    Install Streak Camera:
      Injection/Ramp/Extraction studies…
  SD23 THz detector, 2nd wedged window, mechanical upgrades